## Slide 1: Coagulation Mini-Study
Kyla Bauman

## Slide 2: STAGO ANALYZER
Coagulation analyzer
PT, PTT, Fibrinogen, D-dimer, Heparin Xa
Methodologies Used:
  Viscosity-based (mechanical) detection system
  Immunoturbidometric
  Colorimetric Assay

## Slide 3: STUDY OBJECTIVE
Manufacturer’s Instructions for reagents are timely and unnecessary
  Refrigerated reagent rack
Poses a risk for potentially delaying patient care
  Incident during stat situation
Study Objective: Correlate results obtained using reagent that was allowed to come to room temperature, versus reagent that was loaded cold (directly from the refrigerator).

## Slide 4: ACCEPTANCE CRITERIA
ANALYTE | R Value > | Slope =/- | Medical Director’s TEa
PT | 0.95 | 0.1 | Within 5%
PTT | 0.95 | 0.1 | Within 8%
Fibrinogen | 0.95 | 0.1 | Within 10%
D-Dimer | 0.95 | 0.1 | Within 20%
Anti-Xa | 0.95 | 0.2 | Within 15%

## Slide 5: PROCEDURE OUTLINE
Collect sample pool
  20 Samples per analyte
  Split between each analyzer
  High, Normal, and Low values
2 sets of reagent were prepared
  Room temperature (RT)
  Refrigerated (COLD)
Each analyte was tested using either RT reagent or COLD Reagent
  On each analyzer

## Slide 6: REGRESSION LINE REVIEW
The data was printed and transferred to a spread sheet
Programmed calculations were made
Criteria review to follow
  R value
  Slope
  Total Allowable Error

## Slide 7: Protime Results
Acceptance Criteria PASSED

## Slide 8: PTT Results
Acceptance Criteria PASSED

## Slide 9: Fibrinogen Results
Acceptance Criteria PASSED

## Slide 10: D-Dimer Results
Acceptance Criteria PASSED

## Slide 11: Anti-Xa Results
Acceptance Criteria PASSED

## Slide 12: CONCLUSION
Success!
The necessary criteria was met
No clinically significant effect on patient results
Medical Director approval
Improvement of patient care, especially in stat situations
Study Outcome: Stago reagent can be directly loaded onto the analyzer from the refrigerator.

## Slide 13: What I Learned
Gained more experience with the analyzer
  Navigating modules
  Troubleshooting QC
Stago Reagents
  Reconstitution
  Stability
Hemostasis reference ranges
Introduction into correlation studies in the clinical laboratory

## Slide 14: MULTIPLE CHOICE QUESTION
What methodology does the Stago Coagulation Analyzer use to measure PT, PTT, and Fibrinogen?
A) Chromatographic
B) Nephelometric
C) Immunologic
D) Mechanical/Magnetic